## Aidoc
Artificial Intelligence and Machine Learning (AI/ML)-Enabled Medical Devices
Date of Final Decision | Submission Number | Device | Company | Panel (lead) | Primary Product Code | use in analysis of | Intended to do assist hospital networks and trained radiologists in workflow triage by flagging and communication of: | cases number | Sensitivity | Specificity | Other Performance Data | Other Describe
08/01/2018 | K180647 | BriefCase | Aidoc Medical, Ltd. | Radiology | QAS | non-enhanced head CT and CTPA images | suspected Intracranial Hemorrhage (ICH) | 198 cases from 3 clinical sites (2 US and 1 OUS) | 93.6% (95% CI: 86.6%-97.6%) | 92.3% (95% CI: 85.4%-96.6%) | time-to-notification of all cases: Standard of care time-to-exam-open (72.6 minutes: 95% CI 45.0-100.2) is significantly longer than the parallel time-to-notification of the BriefCase software (4.5 minutes: 95% CI 4.1-4.8). The mean difference of 68.1 minutes (95% CI 40.5-95.7) for these two metrics is statistically significant. | The device does not alter the original medical image and is not intended to be used as a diagnostic device. The results of BriefCase are intended to be used in conjunction with other patient information and based on professional judgment, to assist with triage/prioritization of medical images. Notified clinicians are responsible for viewing full images per the standard of care.
04/15/2019 | K190072 | Briefcase | Aidoc Medical, Ltd. | Radiology | QAS | non-enhanced head CT and CTPA images | Intracranial Hemorrhage (ICH) and Pulmonary Embolism (PE) pathologies | Pilot Study: 2,803 CT pulmonary angiogram (CTPA) exams of patients with suspected PE, from a single center Pivotal Study: 184 cases from 3 clinical sites (2 US and 1 OUS) | Pilot Study: 93.0% (95% CI: 90.2%-95.1%) Pivotal Study: 90.6% (95% CI: 82.2%-95.9%) | Pilot Study: 93.7% (95% CI: 92.7%-94.6%) Pivotal Study: 89.9% (95% CI: 82.2%-95.1%) | Standard of care time-to-exam-open (64.1 minutes: 95% CI 36.6-91.5) is significantly longer than the parallel time-to-notification of the BriefCase software (3.9 minutes: 95% CI: 3.7-4.1). | The device does not alter the original medical image and is not intended to be used as a diagnostic device. The results of BriefCase are intended to be used in conjunction with other patient information and based on their professional judgment, to assist with triage/prioritization of medical images. Notified clinicians are responsible for viewing full images per the standard of care.
05/31/2019 | K190896 | BriefCase | Aidoc Medical, Ltd. | Radiology | QAS | cervical spine CT images | suspected positive findings of linear lucencies in the cervical spine bone in patterns compatible with fractures
12/20/2019 | K192383 | BriefCase | Aidoc Medical, Ltd. | Radiology | QAS | head CTA images | suspected positive findings of Large Vessel Occlusion (LVO) pathologies | head CTA images containing Large Vessel Occlusion (LVO) in 383 cases from 3 US-based clinical sites | 88.8% (95% CI: 81.9%, 93.8 %) | 87.2% (95% CI: 82.5%, 91.1%) | The positive likelihood ratio (PLR) and negative likelihood ratio (NLR) are 6.9 (95% CI: 5.0-9.6) and 0.13 (95% CI: 0.1-0.2), respectively The BriefCase mean time-to-notification for LVO was 3.8 min (95% CI: 3.6-4.0). The time-to-notification for the BriefCase for ICH triage was 4.46 minutes (95% CL: 4.10-4.83). | The device does not alter the original medical image and is not intended to be used as a diagnostic device. The results of BriefCase are intended to be used in conjunction with other patient information and based on their professional judgment, to assist with triage/prioritization of medical images. Notified clinicians are responsible for viewing full images per the standard of care.
06/19/2020 | K193298 | BriefCase | Aidoc Medical, Ltd. | Radiology | QAS | abdominal CT images | suspected positive findings of Intra-abdominal free gas (IFG) pathologies
08/26/2020 | K201020 | BriefCase | Aidoc Medical, Ltd. | Radiology | QAS | contrast enhanced chest CT images (but not dedicated CTPA protocol) CT | suspected positive cases of Incidental Pulmonary Embolism (iPE) pathologies.
04/14/2021 | K202992 | BriefCase, RIB Fractures Triage (RibFx) | Aidoc Medical, Ltd. | Radiology | QFM | chest CT (with or without contrast) | suspected cases of three or more acute Rib fracture (RibFx) pathologies.
12/29/2020 | K203508 | BriefCase | Aidoc Medical, Ltd. | Radiology | QAS | cervical spine CT images | suspected positive findings of linear lucencies in the cervical spine bone in patterns compatible with fractures
03/21/2022 | K213721 | BriefCase | Aidoc Medical, Ltd. | Radiology | QAS | head CT Angio (CTA) images | suspected positive cases of Brain Aneurysm (BA) findings above 5 mm size. | 268 cases, 96 positive cases and 172 negative cases (reports on images with Brain Aneurysm versus without Brain Aneurysm findings) included in the analysis. | 88.5% (95% CI: 80.4%, 94.1%) | 89.5% (95% CI: 84.0%, 93.7%) | NPV was 98.9% (95% CI: 98.3%-99.6%) and PPV was 42.4% (95% CI: 30.0%-51.7%) PLR was 8.46% (95% CI: 5.43%-13.18%) and NLR was 0.13 (95% CI: 0.07%-0.22%) The standard of care time-to-exam-open (89.4 minutes: 95% CI: 56.0-122.7; Median 66.0, IQR 50.7) was significantly longer than the time-to-notification metric of the BriefCase device (4.2 minutes, 95% CI: 3.9-4.5; Median 4.2, IQR 1.8).
04/26/2022 | K213886 | BriefCase | Aidoc Medical, Ltd. | Radiology | QAS | contrast-enhanced chest CTs (not dedicated CTPA protocol) | suspect cases of incidental Pulmonary Embolism (iPE) pathologies
03/14/2022 | K214043 | BriefCase | Aidoc Medical, Ltd. | Radiology | QFM | Chest X-Ray images in adults or transitional adolescents aged 18 and older | suspect positive cases with Pneumothorax (Ptx) findings
10/07/2022 | K220709 | BriefCase | Aidoc Medical, Ltd. | Radiology | QAS | head CTA images | suspected positive findings of complete Large Vessel Occlusion (LVO) - MCA-M1, PCA-P1, ACA-A1, ICA, Basilar; and Medium Vessel Occlusions (MeVO) - MCA-M2, MCA-proximal M3, PCA-P2, PCA-proximal P3, ACA-A2, ACA-proximal A3, and Vertebral-V4 | 342 Vessel Occlusion (VO) cases from 5 US-based clinical sites | 91.3% (95% CI: 83.6%, 96.2%) | 85.6% (95% CI: 80.6%, 89.7%) | NPV was 98.9% (95% CI: 4.66%- 99.4%) and PPV was 41.3% (95% CI: 34.1%- 49.0%) PLR was 6.34% (95% CI: 97.9%- 8.63%) and NLR was 0.10% (95% CI: 0.05%- 0.20%) The BriefCase mean time-to-notification for VO was 2.23 minutes (95% CI: 2.22-2.23) The time-to-notification for the predicate BriefCase LVO was 3.80 minutes (95% CI: 3.60-4.00)
05/17/2022 | K221240 | BriefCase | Aidoc Medical, Ltd. | Radiology | QAS | nonenhanced head CT images | suspected positive findings of Intracranial hemorrhage (ICH) pathologies | head CT images containing intracranial hemorrhage (ICH) in 220 cases from 5 US-based clinical sites. | 96.15% (95% CI: 90.44%, 98.94%) | 94.83% (95% CI: 89.08%, 98.08%) | NPV was 99.29% (95% CI: 98.16% - 99.73%) and PPV was 76.64% (95% CI: 60.06% - 87.74%) PLR was 18.59 (95% CI: 8.52 - 40.56) and NLR was 0.04 (95% CI: 0.02 - 0.11)
06/03/2022 | K221314 | BriefCase | Aidoc Medical, Ltd. | Radiology | QAS | head CTA images | M1 Large Vessel Occlusion (M1-LVO) pathologies (retrospective, blinded, multicenter study) | previously provided in K192383 | same as K192383 | same as K192383 | NPV was 99.5% (95% CI: 99.3%-99.8%) and PPV was 20.1% (95% CI: 13.2%-24.2%) Gender distribution was 39.9% male, and 59.3% female Manufacturer Frequency Distribution of Scanner: Siemens 32, 8%; GE 212, 55%; Canon 73, 19%; Philips 66, 17%
11/18/2022 | K221330 | BriefCase | Aidoc Medical, Ltd. | Radiology | QAS | frontal chest X-ray (CXR) images | suspected positive cases of vertically malpositioned endotracheal tube (ETT) in relation to the carina
08/26/2022 | K222277 | BriefCase | Aidoc Medical, Ltd. | Radiology | QAS | CTPA images | suspected positive findings of Pulmonary Embolism (PE) pathologies
09/28/2022 | K222329 | BriefCase | Aidoc Medical, Ltd. | Radiology | QAS | CT exams with contrast (CTA and CT with contrast) that include the chest | suspected positive cases of aortic dissection (AD) pathology
12/05/2022 | K222692 | BriefCase | Aidoc Medical, Ltd. | Radiology | QFM | chest and abdominal CT images | suspected positive cases of Vertebral Compression Fractures (VCFx) findings
02/01/2023 | K230020 | BriefCase | Aidoc Medical, Ltd. | Radiology | QFM | chest CTs (with or without contrast) images | suspect cases of three or more acute Rib fracture (RibFx) pathologies
11/08/2023 | K230534 | BriefCase-Quantification | Aidoc Medical, Ltd. | Radiology | QIH | abdominal aorta CT exams with contrast | cases that include the abdominal aorta (M-AbdAo)
11/28/2023 | K231631 | BriefCase-Quantification | Aidoc Medical, Ltd. | Radiology | JAK | non-cardiac-gated non-contrast CT (NCCT) images that include the heart | coronary artery calcium scoring
11/13/2023 | K232083 | BriefCase-Quantification | Aidoc Medical, Ltd. | Radiology | QIH | non-contrast head CT (NCCT) images | automatic measurement of brain midline shift | 284 cases from 228 unique patients from 6 US-based clinical sites |  |  | the mean absolute error between the ground truth measurement and algorithm was 0.94 mm (95% CI: 0.74 mm, 1.14 mm) mm
10/30/2023 | K232751 | BriefCase-Triage | Aidoc Medical, Ltd. | Radiology | QAS | CTPA images | suspected positive cases of Central Pulmonary Embolism (Central PE)
05/15/2024 | K241112 | BriefCase-Quantification | Aidoc Medical, Ltd. | Radiology | QIH | abdominal aorta CT exams with contrast | cases that include the abdominal aorta (M-AbdAo)